## Slide 1: DICOM Educational Conference Bangkok, Thailand
3-4 October 2019
DICOMweb™ IN PRACTICE
Brad Genereaux
Medical Imaging Alliance Manager, NVIDIA
Copyright DICOM® 2019 | www.dicomstandard.org | #DICOMConference2019 | #DICOM | @DICOMstandard | @integratorbrad

## Slide 2: Recap: What is DICOMweb?
HTTP-driven services for access to DICOM services, using REST architecture
Incremental capability enhancements for DICOM-enabled systems where necessary
  i.e., image-producing modalities don’t all need to be retrofitted to support DICOMweb – it is leveraged where it is needed
endpoint
resource type
study UID
http://server.org/dicomweb/studies/1.2.345

## Slide 3: Recap: DICOMweb™ Services
Query
QIDO-RS (Query based on ID for DICOM Objects)
DICOM PS3.18 10.6
Retrieve
WADO-RS (Web Access of DICOM Objects)
DICOM PS3.18 10.4
Store
STOW-RS (Store over the web)
DICOM PS3.18 10.5
Tasks
UPS-RS (Worklist Service)
DICOM PS3.18 11
Server Info
DICOM PS3.18 8.9
Capabilities Service

## Slide 4: Recap: Data Structures
<DicomAttribute Tag="00080020" VR="DT" Keyword="StudyDate">
<Value number="1">20130409</value>
</DicomAttribute>
<DicomAttribute Tag="00080061" VR="CS" Keyword="ModalitiesInStudy">
<Value number="1">CT</value>
</DicomAttribute>
<DicomAttribute Tag="00100010" VR="PN" Keyword="PatientName">
<PersonName number="1">
<SingleByte>
<FamilyName>Doe</FamilyName>
<GivenName>John</GivenName>
</PersonName>
</DicomAttribute>
<DicomAttribute Tag="0020000D" VR="UI" Keyword="StudyInstanceUID">
<Value number="1">
1.2.392.200036.9116.2.2.2.1762893313.1029997326.945873
</Value>
</DicomAttribute>
XML
"00080020": {
"vr": "DT", "Value": ["20130409"]
},
"00080061": {
"vr": "CS", "Value": ["CT"]
},
"00100010": {
"vr": "PN", "Value": [ {
"Alphabetic": {
"Family": ["Doe"], "Given": ["John"]
}
}]
},
"0020000D": {
"vr": "UI", "Value": [
"1.2.392.200036.9116.2.2.2.1762893313.1029997326.945873"
]
}
JSON

## Slide 5: Examples of DICOMweb In Practice
Web-enabled workflows
  E.g., web viewers, embedding inside other applications like EMRs
  Non-DICOM-centric applications
Networked workflows
  E.g., remote or foreign storage where firewall or VPN traversal occurs
In IHE Profiles
  Query and Retrieve in WIA (Web Image Access)
  Store (STOW) in WIC (Web Image Capture)
(Your use case here!)

## Slide 6: Sample DICOMweb Tools
The following are just a sample of projects with DICOMweb tools that developers can use to get started:
DICOMweb Servers
  Orthanc (https://www.orthanc-server.com/)
  DCM4CHEE (https://www.dcm4che.org/)
  Google Cloud (https://cloud.google.com/healthcare/docs/how-tos/dicomweb)
Medical Image Viewers
  OHIF (http://ohif.org/)
Public Servers
  SIIM (http://siim.org/hackathon)
  Medical Connections (https://www.dicomserver.co.uk/Pages/Index/RESTful)
Working with REST
  Postman (https://www.getpostman.com/)

## Slide 7: Sample Conformance Statements
The following are just a sample of companies with DICOMweb conformance statements published on their websites:
Agfa: https://global.agfahealthcare.com/main/miscellaneous/interoperability/dicom_connectivity/
DCM4CHEE: https://dcm4chee-arc-cs.readthedocs.io/en/latest/introduction.html
GE: https://www.gehealthcare.com/products/interoperability/dicom
Orthanc: https://www.orthanc-server.com/static.php?page=documentation
Sectra: https://medical.sectra.com/knowledge-center/conformance-statements/
Vital: https://www.vitalimages.com/customer-success-support-program/compliance-documents/

## Slide 8: Query (QIDO-RS) Example
Study | Date | Link
Abdomen CT | Jan. 17, 2014 | http://<url>
Chest CR | Mar. 23, 2015 | http://<url>
What studies do you have for John Doe?
GET http://server.com/studies/?00100010=DOE^JOHN

## Slide 9: QIDO-RS Query Parameters
Key | Value | Description
{attributeID} | {value} | Query matching on supplied value
includefield | {attribute} ¦ all | Include supplied DICOM header values in result
fuzzymatching | true ¦ false | Whether query should use fuzzy matching
limit | {n} | Return only {n} results
offset | {n} | Skip {n} results
E.g.:
https://server/studies​?PatientID=12345678
https://server/studies​?PatientID=12345678​&StudyDate=20180101
https://server/studies​?00100010=DOE*​&00101002.00100020=12345678&limit=25

## Slide 10: Postman: Query

## Slide 11: Retrieve (WADO-RS) Example
Show me images for the Abdomen CT
GET http://server.com/studies/1.2.3/series/4.5.6/

## Slide 12: WADO-RS Sub-Resources and Parameters
Key | Value | Description
annotation | “patient” / “technique” | Add burned-in demographics / procedure details
quality | {n} | Quality of image (lossy factor)
viewport | vw,vh / sx,sy,sw,sh | Width and height, or crop to specific region
window | centre,width,shape | Center of the range of gray scale in the image
Sub-resource | Description
(none) | Retrieve the DICOM objects in the requested format in the Accept header
/rendered | Retrieve a renderable version of the instances (e.g., JPG)
/metadata | Retrieve the DICOM headers in the requested format (JSON or XML)
/bulkdata | Retrieve the bulkdata items in the requested resource
/thumbnail | Retrieve a representative rendered image of the study
Render Options:

## Slide 13: Postman: Retrieve Metadata

## Slide 14: Postman: Retrieve DICOM Part 10

## Slide 15: Store (STOW-RS) Example
Stored
POST http://server.com/studies
Store this image

## Slide 16: Tasks (UPS-RS) Example
Study | Date | Link
Abdomen CT | Nov. 20, 2015 | http://<url>
Abdomen CT | Nov. 20, 2015 | http://<url>
GET http://server.com/workitems/?00741202=BRADLIST
What do I need to do today?

## Slide 17: Capabilities Example
OPTIONS http://server.com/
What can you do for me?
Query
QIDO-RS (Query based on ID for DICOM Objects)
DICOM PS3.18 6.7
Retrieve
WADO-RS (Web Access of DICOM Objects)
DICOM PS3.18 6.5
Store
STOW-RS (Store over the web)
DICOM PS3.18 6.6
Tasks
UPS-RS (Worklist Service)
DICOM PS3.18 6.9
Server Info
DICOM PS3.18 6.8
Capabilities Service

## Slide 18: Using FHIR and DICOMweb for AI Training
Patient | Date
123 | Jan. 17, 2014
456 | Mar. 23, 2015
Get me all patients who have had pneumonia.
GET https://server.com/fhir/Patient?condition.code=233604007
Get me the CR images for these studies.
GET http://server.com/studies/?00100010=DOE^JOHN
Study | Type
1.2.3 | X-Ray Chest
4.5.6 | CT Head
GET https://dicom.com/studies/1.2.3
Get me all studies for these patients.

## Slide 19: DICOMweb™ Cheatsheet
Available at: https://www.dicomstandard.org/dicomweb-cheatsheet/
Available at: https://www.dicomstandard.org/dicomweb-cheatsheet/

## Slide 20: Questions? Thank you!
Presenter
Brad Genereaux, Medical Imaging Alliance Manager - NVIDIA
bgenereaux@nvidia.com
@IntegratorBrad
https://linkedin.com/in/integratorbrad